**AVI NITRI SKIN PRO PRODUCT SPECIFICATIONS**

|  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| **Name of Product** | |  | : 3 Mil Powder Free Nitrile Exam Gloves | | | | | |  |
| **Name of Manufacturer**  Innovative Gloves Co. Ltd. Hatyai, Thailand | | | | | | | | |  |
| **Base Material** | |  | : Nitrile | |  |  |  |  |  |
| **Characteristics** | |  | : Non Sterile, Ambidextrous Beaded Cuff | | | | | |  |
| **Available Sizes** | |  | : X- Small,Small, Medium, Large, X-Large, XX-Large | | | | | | |
| **Surface** | |  | : Finger Textured. | |  |  |  |  |  |
| **Surface Treatment** | | | : Polymer Coating | |  |  |  |  |  |
| **Colour** | |  | : White/Blue/Black | |  |  |  |  |  |
| **Physical Specifications:** | | |  |  |  |  |  |  |  |
|  |  | Length | | Thickness(mm) | | |  |  |  |
| Size |  | SINGLE WALL | | |  |  | Width (mm) |
|  | (mm) |  |  |  |
|  |  |  | Tip | Palm |  | Cuff |  |  |
|  |  |  |  |  |  |  |
| X-Small |  | 240±5 | | 0.09±0.01 | 0.06±0.01 |  | 0.05±0.01 |  | 75±5 |
| Small |  | 240±5 | | 0.09±0.01 | 0.06±0.01 |  | 0.05±0.01 |  | 85±5 |
| Medium |  | 240±5 | | 0.09±0.01 | 0.06±0.01 |  | 0.05±0.01 |  | 95±5 |
| Large |  | 240±5 | | 0.09±0.01 | 0.06±0.01 |  | 0.05±0.01 |  | 105±5 |
| X-Large |  | 240±5 | | 0.09±0.01 | 0.06±0.01 |  | 0.05±0.01 |  | 115±5 |
| XX-Large |  | 240±5 | | 0.09±0.01 | 0.06±0.01 |  | 0.05±0.01 |  | 120±3 |

**A picture containing handwear, clothing

Description generated with very high confidence**

|  |  |  |
| --- | --- | --- |
| **Tensile(MPA)** | **Elongation(%)** | **Force At Break(N)** |
| **22 Mpa min.** | **400min.** | **≥ 6 N** |

**Tensile (ASTM D 6319 & EN 455-2)**

**Before Aging**

**After Accelerated Aging**

|  |  |
| --- | --- |
| **Tensile(MPA)** | **Elongation(%)** |
| **22 Mpa min.** | **500 min.** |

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Total Residue Content** | | | : < 2.0 Milligrams per glove | |
| **Packaging (Normal)** | |  | : 100±2 pieces in one Dispenser Box | | |
|  |  |  | 10 Dispenser Box in one Carton | | |
| **Labeling on Dispenser (Normal)** | | | : Product Name, Size, Lot no. Quantity, | | |
|  |  |  | Manufacturer’s address, Warning about Latex Allergy | | |
| **Labeling on Carton (Normal)** | | | : Product Name, Size, Lot no. Quantity, Country of | | |
|  |  |  | Origin |  |

**Container Loadability** (No MOQ in manufacturer owned brand in standard Packaging)

20FCL=2150 for 100x10 Packing,1360 for 200x10 Packing